## Slide 1: Using the EyeLink 1000 Plus UCLA Cognitive Area
Ringo Huang
Lab Meeting
03/08/2022

## Slide 2: Overview
Experiment design for pupillometry
  How to correct for gaze position
  How to luminance norm
Programming your experiment
Running your study
Preprocessing your data

## Slide 3: Materials
Materials for this demo can be found here:
https://github.com/RingoHHuang/ucla-eyelink-demo
Two templates:
picture.py (pictures in images folder)
sound.py (audio files in sounds folder)
Helper scripts:
shine_example.m (luminance matching)
edf2struct.m (edf output to matlab data structure)

## Slide 4: Experiment Design for Pupillometry

## Slide 5: How do experimenters adjust for changes to pupil size based on gaze location?
Pupil foreshortening error (PFE) – the pupil image is a 2D projection whose area decreases as the eye rotates away from the camera
Set-up to minimize PFE:
Place camera as close to the line-of-sight between the pupil and stimuli as possible
Use pupil data from the eye that is closest to the lens of the camera
Place screen as far back from the participant as possible
In general, PFE increases as increases

## Slide 6: How do we reduce Pupil Foreshortening Error in our experiments?
Do nothing
Reduces statistical power
May introduce systematic biases leading to Type 1 errors
Control fixation
Eliminates PFE
No free viewing
Constrain viewing area
Permits some free viewing
Introduces some PFE
Smaller viewing area reduces the ability to resolve small changes in gaze
Constrain viewing along a fixed arc
Controls the PFE between trials
Only useful in special cases
+

## Slide 7: How do we correct for Pupil Foreshortening Error?
PFE correction (Hayes and Petrov, 2015)
PFE =
Raw Pupil Diameter,
Gaze X, Gaze Y
Pupil Foreshortening Error
Corrected Pupil Diameter
(Raw PD/PFE)

## Slide 8: Stimuli Considerations
Auditory Tones
can remove glasses
minimizes noise in pupil data
Isoluminant
Grayscale Images
- use SHINE toolbox to luminance norm the set of images
Isoluminant
Color Images
- can apply SHINE luminance matching to each RGB color channel
Recommend having a central fixation cross during all interstimulus intervals

## Slide 9: Luminance matching with SHINE
Do all other image processing steps prior to luminance matching
Convert to grayscale, then find average luminance of the set of images
lumMatch
SHINE toolbox download: http://www.mapageweb.umontreal.ca/gosselif/shine/
shine_example.m demonstrates this procedure

## Slide 10: Programming your experiment (Using PsychoPy)

## Slide 11: SR support forum
The SR support forum has resources and guides for integrating the EyeLink API with your experiment
Sign up for an account: https://www.sr-support.com
The Software Development Kit comes with documentations and example experiments for the most common stimuli presentation languages, including E-Prime, MATLAB, Python, and more.
This demo will focus on PsychoPy

## Slide 12: Hardware Overview
EyeLink on/off switch
Remember to switch OFF after use!
BLUE ethernet cable connects to experiment display PC
HDMI cable connects to experiment display PC
Host PC controls the eye-tracker
Shut down after use!
Head Rest
Stimuli Monitor
EyeLink
Display PC
Host PC

## Slide 13: A PsychoPy demonstration
A barebones PsychoPy template based on the picture.py example
  Trimmed off 1/3rd of the picture.py script
https://github.com/RingoHHuang/ucla-eyelink-demo
  picture.py – displays all images in “images” folder
  sound.py – plays all audio files in “sounds” folder

## Slide 14: Set up your display PC
Configure your Display PC’s network connection to communicate with the EyeLink Host PC
  https://www.sr-support.com/showthread.php?tid=281
Install the latest version of the EyeLink SDK for your OS
  https://www.sr-support.com/showthread.php?tid=13
3. Install and configure PyLink
  https://www.sr-support.com/thread-48.html
  PyLink is the Python wrapper for the EyeLink API
The display PC is the computer you’ll use to control stimuli presentation. You can use your personal laptop or the Mac Mini in the eye-tracking room.

## Slide 15: Organize experiment directory
One option for organizing your experiment directory:
project_folder
  │ experiment_script.py
  │ EyeLinkCoreGraphicsPsychoPy.py
  │-- results
  │ │ participant_id1_datetime.edf │ │ participant_id2_datetime.edf
  │ │ participant_id3_datetime.edf
  │-- images
  │ │ image1.jpg │ │ image2.jpg
  │-- sounds
  │ │ audio1.wav │ │ audio2.wav
Some convenience functions are in EyeLinkCoreGraphicsPsychoPy.py, so make sure they are accessible

## Slide 16: Start up dialog
# Retrieve participant id, which will be used in the edf filename
# The EDF data filename should not exceed 8 alphanumeric characters
# use ONLY number 0-9, letters, & _ (underscore) in the filename
participant_id = 'TEST'
# Prompt user to specify an EDF data filename
# before we open a fullscreen window
dlg_title = 'Enter participant ID'
dlg_prompt = 'Please enter a ID with 8 or fewer characters\n' + \
'[letters, numbers, and underscore].'
# loop until we get a valid filename
while True:
dlg = gui.Dlg(dlg_title)
dlg.addText(dlg_prompt)
dlg.addField('participant id:', participant_id)
# show dialog and wait for OK or Cancel
ok_data = dlg.show()
if dlg.OK: # if ok_data is not None
print('EDF data filename: {}'.format(ok_data[0]))
else:
print('user cancelled')
core.quit()
sys.exit()
# get the string entered by the experimenter
# strip trailing characters
participant_id = dlg.data[0]
edf_fname = participant_id.rstrip()
# check if the filename is valid (length <= 8 & no special char)
allowed_char = ascii_letters + digits + '_'
if not all([c in allowed_char for c in edf_fname]):
print('ERROR: Invalid EDF filename')
elif len(edf_fname) > 8:
print('ERROR: EDF filename should not exceed 8 characters')
else:
break

## Slide 17: Create edf output filename
# Set up a folder to store the EDF data files and the associated resources
# e.g., files defining the interest areas used in each trial
results_folder = 'results'
if not os.path.exists(results_folder):
os.makedirs(results_folder)
# We download EDF data file from the EyeLink Host PC to the local hard
# drive at the end of each testing session, here we rename the EDF to
# include session start date/time
time_str = time.strftime("_%Y_%m_%d_%H_%M", time.localtime())
session_identifier = edf_fname + time_str
The edf output is located in the “results” subfolder of the project
The edf filename is derived from the participant ID and the time of recording

## Slide 18: Set up connections
# Step 1: Connect to the EyeLink Host PC
#
# The Host IP address, by default, is "100.1.1.1".
# the "el_tracker" objected created here can be accessed through the Pylink
try:
el_tracker = pylink.EyeLink("100.1.1.1")
except RuntimeError as error:
print('ERROR:', error)
core.quit()
sys.exit()
Creates a new pylink object
Establishes a link between EyeLink Host PC and the Python script
# Step 2: Open an EDF data file on the Host PC
edf_file = edf_fname + ".EDF"
try:
el_tracker.openDataFile(edf_file)
except RuntimeError as err:
print('ERROR:', err)
# close the link if we have one open
if el_tracker.isConnected():
el_tracker.close()
core.quit()
sys.exit()
Opens a new edf data file for writing on the Host PC

## Slide 19: Configure the recording
Specify the variables that you want to record (file_event_flags) and the variables that you may want to access in your Python script (link_event_flags ??)
# Step 3: Configure the tracker
#
# Put the tracker in offline mode before we change tracking parameters
el_tracker.setOfflineMode()
# Get the software version: 1-EyeLink I, 2-EyeLink II, 3/4-EyeLink 1000,
# 5-EyeLink 1000 Plus, 6-Portable DUO
vstr = el_tracker.getTrackerVersionString()
eyelink_ver = int(vstr.split()[-1].split('.')[0])
# print out some version info in the shell
print('Running experiment on %s, version %d' % (vstr, eyelink_ver))
# File and Link data control
# what eye events to save in the EDF file, include everything by default
file_event_flags = 'LEFT,RIGHT,FIXATION,SACCADE,BLINK,MESSAGE,BUTTON,INPUT'
# what eye events to make available over the link, include everything by default
link_event_flags = 'LEFT,RIGHT,FIXATION,SACCADE,BLINK,BUTTON,FIXUPDATE,INPUT'
# what sample data to save in the EDF data file and to make available
# over the link, include the 'HTARGET' flag to save head target sticker
# data for supported eye trackers
if eyelink_ver > 3:
file_sample_flags = 'LEFT,RIGHT,GAZE,HREF,RAW,AREA,HTARGET,GAZERES,BUTTON,STATUS,INPUT'
link_sample_flags = 'LEFT,RIGHT,GAZE,GAZERES,AREA,HTARGET,STATUS,INPUT'
else:
file_sample_flags = 'LEFT,RIGHT,GAZE,HREF,RAW,AREA,GAZERES,BUTTON,STATUS,INPUT'
link_sample_flags = 'LEFT,RIGHT,GAZE,GAZERES,AREA,STATUS,INPUT'
el_tracker.sendCommand("file_event_filter = %s" % file_event_flags)
el_tracker.sendCommand("file_sample_data = %s" % file_sample_flags)
el_tracker.sendCommand("link_event_filter = %s" % link_event_flags)
el_tracker.sendCommand("link_sample_data = %s" % link_sample_flags)

## Slide 20: Configure the calibration
Here, you set up the calibration/validation environment
Optional configurations are not included here
# Step 4: set up a graphics environment for calibration
#
# Choose a calibration type, H3, HV3, HV5, HV13 (HV = horizontal/vertical),
el_tracker.sendCommand("calibration_type = HV9")
# Open a window, be sure to specify monitor parameters
mon = monitors.Monitor('myMonitor', width=53.0, distance=70.0)
win = visual.Window(fullscr=True,
monitor=mon,
screen=1,
winType='pyglet',
units='pix')
# get the native screen resolution used by PsychoPy
scn_width, scn_height = win.size
# Pass the display pixel coordinates (left, top, right, bottom) to the tracker
# see the EyeLink Installation Guide, "Customizing Screen Settings"
el_coords = "screen_pixel_coords = 0 0 %d %d" % (scn_width - 1, scn_height - 1)
el_tracker.sendCommand(el_coords)
# Write a DISPLAY_COORDS message to the EDF file
# Data Viewer needs this piece of info for proper visualization, see Data
# Viewer User Manual, "Protocol for EyeLink Data to Viewer Integration"
dv_coords = "DISPLAY_COORDS 0 0 %d %d" % (scn_width - 1, scn_height - 1)
el_tracker.sendMessage(dv_coords)
# Configure a graphics environment (genv) for tracker calibration
genv = EyeLinkCoreGraphicsPsychoPy(el_tracker, win)
print(genv) # print out the version number of the CoreGraphics library
# Set background and foreground colors for the calibration target
# in PsychoPy, (-1, -1, -1)=black, (1, 1, 1)=white, (0, 0, 0)=mid-gray
foreground_color = (-1, -1, -1)
background_color = win.color
genv.setCalibrationColors(foreground_color, background_color)
# Request Pylink to use the PsychoPy window we opened above for calibration
pylink.openGraphicsEx(genv)

## Slide 21: Optional calibration specifications
If you want to specify any of these optional specifications, they would go before the last line of code of the previous slide
Refer to manual for details about these options
# Set up the calibration target
#
# The target could be a "circle" (default), a "picture", a "movie" clip,
# or a rotating "spiral". To configure the type of calibration target, set
# genv.setTargetType to "circle", "picture", "movie", or "spiral", e.g.,
# genv.setTargetType('picture')
#
# Use gen.setPictureTarget() to set a "picture" target
# genv.setPictureTarget(os.path.join('images', 'fixTarget.bmp'))
#
# Use genv.setMovieTarget() to set a "movie" target
# genv.setMovieTarget(os.path.join('videos', 'calibVid.mov'))
# Use a picture as the calibration target
genv.setTargetType('picture')
genv.setPictureTarget(os.path.join('images', 'fixTarget.bmp'))
# Configure the size of the calibration target (in pixels)
# this option applies only to "circle" and "spiral" targets
# genv.setTargetSize(24)
# Beeps to play during calibration, validation and drift correction
# parameters: target, good, error
# target -- sound to play when target moves
# good -- sound to play on successful operation
# error -- sound to play on failure or interruption
# Each parameter could be ''--default sound, 'off'--no sound, or a wav file
genv.setCalibrationSounds('', '', '')
OPTIONAL

## Slide 22: Define helper functions
def terminate_task():
""" Terminate the task gracefully and retrieve the EDF data file
file_to_retrieve: The EDF on the Host that we would like to download
win: the current window used by the experimental script
"""
el_tracker = pylink.getEYELINK()
if el_tracker.isConnected():
# Terminate the current trial first if the task terminated prematurely
error = el_tracker.isRecording()
if error == pylink.TRIAL_OK:
abort_trial()
# Put tracker in Offline mode
el_tracker.setOfflineMode()
# Clear the Host PC screen and wait for 500 ms
el_tracker.sendCommand('clear_screen 0')
pylink.msecDelay(500)
# Close the edf data file on the Host
el_tracker.closeDataFile()
# Show a file transfer message on the screen
msg = 'EDF data is transferring from EyeLink Host PC...'
show_msg(win, msg, wait_for_keypress=False)
# Download the EDF data file from the Host PC to a local data folder
# parameters: source_file_on_the_host, destination_file_on_local_drive
local_edf = os.path.join(session_folder, session_identifier + '.EDF')
try:
el_tracker.receiveDataFile(edf_file, local_edf)
except RuntimeError as error:
print('ERROR:', error)
# Close the link to the tracker.
el_tracker.close()
# close the PsychoPy window
win.close()
# quit PsychoPy
core.quit()
sys.exit()
def clear_screen(win):
""" clear up the PsychoPy window"""
win.fillColor = genv.getBackgroundColor()
win.flip()
def show_msg(win, text, wait_for_keypress=True):
""" Show task instructions on screen"""
msg = visual.TextStim(win, text,
color=genv.getForegroundColor(),
wrapWidth=scn_width/2)
clear_screen(win)
msg.draw()
win.flip()
# wait indefinitely, terminates upon any key press
if wait_for_keypress:
event.waitKeys()
clear_screen(win)
def abort_trial():
"""Ends recording """
el_tracker = pylink.getEYELINK()
# Stop recording
if el_tracker.isRecording():
# add 100 ms to catch final trial events
pylink.pumpDelay(100)
el_tracker.stopRecording()
# clear the screen
clear_screen(win)
# Send a message to clear the Data Viewer screen
bgcolor_RGB = (116, 116, 116)
el_tracker.sendMessage('!V CLEAR %d %d %d' % bgcolor_RGB)
# send a message to mark trial end
el_tracker.sendMessage('TRIAL_RESULT %d' % pylink.TRIAL_ERROR)
return pylink.TRIAL_ERROR

## Slide 23: Trial routine function (PyLink Calls)
Only showing pylink-related code
# Start recording
# arguments: sample_to_file, events_to_file, sample_over_link,
# event_over_link (1-yes, 0-no)
try:
el_tracker.startRecording(1, 1, 1, 1)
except RuntimeError as error:
print("ERROR:", error)
abort_trial()
return pylink.TRIAL_ERROR
# Allocate some time for the tracker to cache some samples
pylink.pumpDelay(100)
# get a reference to the currently active EyeLink connection
el_tracker = pylink.getEYELINK()
# put the tracker in the offline mode first
el_tracker.setOfflineMode()
# clear the host screen before we draw the backdrop
el_tracker.sendCommand('clear_screen 0')
# send a "TRIALID" message to mark the start of a trial, see Data
# Viewer User Manual, "Protocol for EyeLink Data to Viewer Integration"
el_tracker.sendMessage('TRIALID %d' % trial_index)
# record_status_message : show some info on the Host PC
# here we show how many trial has been tested
status_msg = 'TRIAL number %d' % trial_index
el_tracker.sendCommand("record_status_message '%s'" % status_msg)
# put tracker in idle/offline mode before recording
el_tracker.setOfflineMode()
At the start of this trial, retrieve the currently active EyeLink connection.
Then, send some messages to the Host PC.
Start recording at the beginning of each trial.

## Slide 24: Trial routine function (PyLink Calls)
# clear the screen
clear_screen(win)
el_tracker.sendMessage('blank_screen_%d' % trial_index)
# show the image, and log a message to mark the onset of the image
clear_screen(win)
img.draw()
win.flip()
el_tracker.sendMessage('image_onset_%d' % trial_index)
img_onset_time = core.getTime() # record the image onset time
Messages are useful for aligning the timestamps of the eye-tracking samples to task events.
This message marks the onset of an image and its trial #
Log messages in the edf file to mark task events.
This message marks the offset of an image and its trial #

## Slide 25: Trial routine function (PyLink Calls)
# record trial variables to the EDF data file, for details, see Data
# Viewer User Manual, "Protocol for EyeLink Data to Viewer Integration"
el_tracker.sendMessage('!V TRIAL_VAR image %s' % pic)
el_tracker.sendMessage('!V TRIAL_VAR RT %d' % RT)
# send a 'TRIAL_RESULT' message to mark the end of trial, see Data
# Viewer User Manual, "Protocol for EyeLink Data to Viewer Integration"
el_tracker.sendMessage('TRIAL_RESULT %d' % pylink.TRIAL_OK)
# stop recording; add 100 msec to catch final events before stopping
pylink.pumpDelay(100)
el_tracker.stopRecording()
Note: In most cases, I would not start/stop recording for each trial. It’s often useful to record pupil data during interstimulus intervals.
These messages are only necessary if you want to use EyeLink Data Viewer for analyses (I think)
Stop recording at the end of the trial.

## Slide 26: Run doTrackerSetup()
# Step 5: Set up the camera and calibrate the tracker
# Show the task instructions
task_msg = 'In the task, you may press the SPACEBAR to end a trial\n' + \
'\nPress Ctrl-C to if you need to quit the task early\n'
task_msg = task_msg + '\nNow, press ENTER twice to calibrate tracker'
show_msg(win, task_msg)
try:
el_tracker.doTrackerSetup()
except RuntimeError as err:
print('ERROR:', err)
el_tracker.exitCalibration()
This convenience function pulls up the calibration/validation sequences as well as lets you adjust the eye-tracker settings.

## Slide 27: Run experiment trials
# Step 7: disconnect, download the EDF file, then terminate the task
terminate_task()
# Step 6: Run the experimental trials, index all the trials
# Grab all images from "images folder"
trials = os.listdir('images')
# randomize the trial list
random.shuffle(trials)
trial_index = 1
for trial_pars in trials:
run_trial(trial_pars, trial_index)
trial_index += 1
This section is your normal PsychoPy experiment script.
After completing the experiment, call the terminate_task() helper function to clean up the EyeLink connections.

## Slide 28: Drift Check (Optional)
# drift check
# we recommend drift-check at the beginning of each trial
# the doDriftCorrect() function requires target position in integers
# the last two arguments:
# draw_target (1-default, 0-draw the target then call doDriftCorrect)
# allow_setup (1-press ESCAPE to recalibrate, 0-not allowed)
#
while not 1:
# terminate the task if no longer connected to the tracker or
# user pressed Ctrl-C to terminate the task
if (not el_tracker.isConnected()) or el_tracker.breakPressed():
terminate_task()
return pylink.ABORT_EXPT
# drift-check and re-do camera setup if ESCAPE is pressed
try:
error = el_tracker.doDriftCorrect(int(scn_width/2.0),
int(scn_height/2.0), 1, 1)
# break following a success drift-check
if error is not pylink.ESC_KEY:
break
except:
pass
If you are doing gaze – it might be worthwhile to include a drift check every so often.
If drift check fails, you can recalibrate.

## Slide 29: Running the session

## Slide 30: Light Meter
You may want to control the luminance level of the eye-tracking room between sessions
  100 lux (?)
iOS App: Light Meter LM-3000
Make your own light diffuser with a strip of paper and tape
Hold RIGHT to DIM
Hold LEFT to BRIGHTEN

## Slide 31: Before the participant arrives
Check the luminance of the room
  Turn on lights (push left button once)
  Close door and make sure all other light sources are controlled (i.e., the stimulus presentation screen is on, host PC monitor is on, etc.)
  Cover your camera lens with the strip of paper
  Align your phone so that it’s just above the light switch panel
  Brighten/dim the lights until your light meter app reads 100 lx (or whatever value you decide)
Hold RIGHT to DIM
Hold LEFT to BRIGHTEN
Turn on Host PC and Display PC
Attach the blue ethernet cable and the HDMI cable to your Display PC (use the USB-C dongle if necessary)

## Slide 32: Start up dialog
Showing what happens when you enter more than 8 characters in the prompt

## Slide 33: doTrackerSetup() – Set up
el_tracker.doTrackerSetup()
Toggle between camera views using L/R arrow keys.
Set-up:
Search boundaries
CR threshold
Pupil threshold
You can also do the same adjustments on the Host PC

## Slide 34: doTrackerSetup() – Calibration/Validation
el_tracker.doTrackerSetup()
Press C to calibrate.
When the circle appears, press spacebar
Automatically advance once fixation is established
Press spacebar to accept calibration
Press V to validate
Note that calibration/validation can also be controlled on the Host PC
Notes: The video capture is bad

## Slide 35: Start Experiment
Press O to start the main experiment.
This completes doTrackerSetup() convenience function and proceeds to the rest of the script.

## Slide 36: Data output
.edf files are saved in your results folder

## Slide 37: Data Processing

## Slide 38: Removing artifacts from your pupil data
Raw Pupil Data
Matlab data structure (.mat)
ET-remove-artifacts
Preprocessed
Pupil Data
ET-Remove-Artifacts Pipeline (github.com/EmotionCognitionLab/ET-remove-artifacts)

## Slide 39: For UCLA Cognitive Area EyeLink System:
edf file
Matlab data structure
Data structure compatible with ET-remove-artifacts
edfmex
edf2struct.m
ET-remove-artifacts
Cleaned Pupil Data
Calculate pupil dilation, aggregate across trials, create timecourse plots, stats, etc.

## Slide 40: Convert edf file to Matlab data structure
Download edfmex from SR support forum and add it to your path
edf2struct.m (included in the eyelink_example project) converts edf to a data structure that is compatible with ET-remove-artifacts
  Reads all .edf files in “results” sub-folder
  Outputs as .mat files in “pupil_data” sub-folder
  Two outputs per sub; one per eye
Snippet from edf2struct.m

## Slide 41: ET-remove-artifacts
github.com/EmotionCognitionLab/ET-remove-artifacts
See the documentation on the Github repository for details not covered here.

## Slide 42: ET-remove-artifacts (algorithm)
The blink detection algorithm:
Derive the velocity plot of the raw pupil data
  In terms of velocity, blinks are characterized as a large negative velocity followed by a large positive velocity
2. Detect the onset and offset of blinks
3. Linearly interpolate across the blink, using the onset and offset pupil sample as endpoints

## Slide 43: ET-remove-artifacts (default settings)
When you load in your data, the artifact removal algorithm is automatically applied with the default settings
Detects the sampling rate of your data
Specifies a filter order for the differentiator filter based on the sample rate
Detects blinks based on the blink detection algorithm
Also detects invalid samples which are samples that fall below 30% of the range of all samples
Imputes artifacts longer than 2s with NaNs

## Slide 44: ET-remove-artifacts (differentiator filter)
The first 3 parameters control the filtered derivative used to get the velocity plot.
A larger filter order and smaller frequency band (i.e., difference between passband and stopband) makes the velocity plot smoother.
A smaller filter order and larger frequency band makes the blink profiles more prominent.

## Slide 45: ET-remove-artifacts (manual)
If there are artifacts that weren’t properly taken care of by the algorithm, you can manually specify a start and end point to linearly interpolate over:

## Slide 46: Example of pupil dilation to both mental effort and contextual/perceptual shift

## Slide 47: Thank you!